**Infusion pumps**

technical features

* Volumetric or peristaltic operation
* Infusing from bag, bottle
* program for succession drug infusion
* Availability of a wide range of special infusion sets (for blood, for photosensitive drugs, for oncological drugs, with / without injection point)
* wide setting range (from 1 to 999 ml / h)
* Total infusion volume settable between 1 - 9999 ml
* Large and bright display for easy reading of all parameters and any alarms
* Accuracy ± 5%
* Alarms for Occlusion, End of infusion, Low battery, Air on line
* Rechargeable battery (autonomy for at least 3 h with an infusion rate of 125 ml / h)
* preferably compliant with the Italian medical devices legislation in force in Italy applied to medical devices.
* Supply: 220 V, 50 Hz
* IEC 60601-1 class;

infusion sets for volumetric infusion characteristics

* EVA or PVC latex free material
* Distal luer lock fitting
* Transparent, non-rigid drip chamber
* Length ≥ 2 meters
* anti-free-flow device